# M30 Sensor Configuration Reference

Platform: M30 VerG  |  Last updated: 2026-05-14

This document describes every parameter that changes when you swap sensors on the M30 platform. It covers PID and electrochemical (EC) sensor types and provides ready-to-paste code for each profile.

## 1  Quick-Swap Checklist

| Step | File | What to change |
| --- | --- | --- |
| 1 | SensorConfig.h | Uncomment the correct #include line |
| 2 | .ino — gas list | g_gasManager.addGas(...) calls |
| 3 | .ino — LIBRARY menu | GasMenuItem(...) calls |
| 4 | .ino — Range menu | RangeMenuItem(...) calls |
| 5 | .ino — Alarm menu | m_alarmArray values |
| 6 | .ino — PEL/Low alarm | Loop values |
| 7 | .ino — Calvalue menu | Loop or array values |
| 8 | inc/RangeSet.h | m_rangeArray values |
| 9 | inc/AlarmSet.h | m_alarmArray values |
| 10 | inc/LowAlarmSet.h | m_lowalarmArray values |
| 11 | inc/CalvalueSet.h | m_calvalueArray values |

Steps 1-7 are in the .ino.  Steps 8-11 update the backing arrays (must match the menu entries in the .ino).

The following change automatically from SensorConfig.h (no manual edit needed):

GasManager constructor (intercept, slope x4 slots)

Bluetooth device name

ADS1115 gain setting

## 2  Parameter Reference

### 2.1  GasManager Constructor

GasManager g_gasManager(

intercept,   slope,    // Gas slot 1 - primary range

zero,        calgas,   // working vars (leave 0, 0)

intercept2,  slope2,   // Gas slot 2

intercept3,  slope3,   // Gas slot 3

intercept4,  slope4    // Gas slot 4

);

| Parameter | Meaning | Typical range |
| --- | --- | --- |
| intercept | Sensor voltage at zero gas (V) | -5 V to +5 V |
| slope | Voltage change per ppm (V/ppm) | 0.001 to 10 |

Rule: slope must never be 0 - it is the denominator in calculateSLM(). Set to 1.0 as the safe boot default; field calibration will overwrite it in EEPROM.

### 2.2  ADS1115 Gain

| Constant | Full-scale | Resolution | Use when |
| --- | --- | --- | --- |
| GAIN_TWOTHIRDS | +/-6.144 V | 0.1875 mV/LSB | Very wide range, low-noise environment |
| GAIN_ONE | +/-4.096 V | 0.1250 mV/LSB | PID sensors (default M30) |
| GAIN_TWO | +/-2.048 V | 0.0625 mV/LSB | EC sensors, amp output 1-2 V |
| GAIN_FOUR | +/-1.024 V | 0.0313 mV/LSB | EC sensors, small output < 1 V |
| GAIN_EIGHT | +/-0.512 V | 0.0156 mV/LSB | Very high-gain TIA, uV signals |

Warning: if the sensor output exceeds the selected full-scale, readings will saturate. Always verify the TIA output voltage at full-scale concentration before choosing gain.

### 2.3  Range Menu  (RangeSet.h)

The m_rangeArray values are the full-scale limits. The run screen displays "xxx" when sensorValue > range. Must match the RangeMenuItem(...) entries in the .ino.

### 2.4  Alarm (High) Menu  (AlarmSet.h)

First entry must be 0 (= Off). Values are the concentration threshold (ppm) above which the buzzer activates and the display flashes ALARM.

### 2.5  PEL / Low Alarm Menu  (LowAlarmSet.h)

First entry must be 0 (= Off). Intended for sustained low-level exposure limits (OSHA PEL or NIOSH REL). Alarm activates when sensorValue > lowalarm.

### 2.6  Calvalue Menu  (CalvalueSet.h)

These are the available certified reference gas concentrations for span calibration. They should match the standard cylinders you have on hand. The selected value is read by GasManager::calibrate2() to compute slope.

## 3  Sensor Profiles

### 3.1  PID - Dual Range  100 / 1000 ppm  (current M30)

| Parameter | Value |
| --- | --- |
| Profile file | inc/sensors/PID_M10_100_1000.h |
| BT name | ESP32 PID M10 |
| ADS gain | GAIN_ONE |
| Intercept (slot 1) | -2.054456771 |
| Slope (slot 1) | 1.0 (factory default; EEPROM overwrites after cal) |
| Intercept (slot 2) | 0.0 |
| Slope (slot 2) | 1.0 |
| Gas library | "100", "1000" |
| Range options | 5000 ppm, 10000 ppm |
| Calvalue options | 10-100 ppm (step 10), 200-500 ppm (step 100) |
| High alarm | 0-1000 ppm (0=Off, step 10 to 500, step 100 to 1000) |
| PEL alarm | 0-60 ppm (step 1) |

### 3.2  EC - CO (Carbon Monoxide)  0-300 ppm

| Parameter | Value |
| --- | --- |
| Profile file | inc/sensors/EC_CO_300ppm.h |
| BT name | ESP32 EC-CO |
| ADS gain | GAIN_FOUR |
| Intercept (slot 1) | Calibrate in field (boot default: 0.0) |
| Slope (slot 1) | Calibrate in field (boot default: 1.0) |
| Gas library | "CO" |
| Range options | 100 ppm, 300 ppm |
| Calvalue options | 10, 25, 50, 100 ppm |
| High alarm default | 50 ppm (OSHA PEL) |
| PEL alarm default | 35 ppm (NIOSH REL) |
| OSHA PEL | 50 ppm TWA (8 h) |
| NIOSH REL | 35 ppm TWA; STEL 200 ppm |
| IDLH | 1200 ppm |

Ready-to-paste arrays for EC CO:

// inc/RangeSet.h

std::array<int, 2> m_rangeArray{{ 100, 300 }};

// inc/AlarmSet.h

std::array<int, 31> m_alarmArray{{

0,

10, 20, 30, 40, 50, 60, 70, 80, 90, 100,

110, 120, 130, 140, 150, 160, 170, 180, 190, 200,

210, 220, 230, 240, 250, 260, 270, 280, 290, 300

}};

// inc/LowAlarmSet.h

std::array<int, 11> m_lowalarmArray{{

0, 5, 10, 15, 20, 25, 30, 35, 40, 45, 50

}};

// inc/CalvalueSet.h

std::array<int, 4> m_calvalueArray{{ 10, 25, 50, 100 }};

### 3.3  EC - H2S (Hydrogen Sulfide)  0-50 ppm

| Parameter | Value |
| --- | --- |
| Profile file | inc/sensors/EC_H2S_50ppm.h |
| BT name | ESP32 EC-H2S |
| ADS gain | GAIN_FOUR |
| Intercept (slot 1) | Calibrate in field (boot default: 0.0) |
| Slope (slot 1) | Calibrate in field (boot default: 1.0) |
| Gas library | "H2S" |
| Range options | 10 ppm, 50 ppm |
| Calvalue options | 1, 2, 5, 10, 25 ppm |
| High alarm default | 10 ppm (OSHA ceiling warning) |
| PEL alarm default | 1 ppm (ACGIH TLV-TWA) |
| OSHA PEL | 20 ppm ceiling |
| NIOSH REL | 1 ppm ceiling (10 min) |
| IDLH | 50 ppm |

Ready-to-paste arrays for EC H2S:

// inc/RangeSet.h

std::array<int, 2> m_rangeArray{{ 10, 50 }};

// inc/AlarmSet.h

std::array<int, 51> m_alarmArray{{

0,  1,  2,  3,  4,  5,  6,  7,  8,  9, 10,

11, 12, 13, 14, 15, 16, 17, 18, 19, 20,

21, 22, 23, 24, 25, 26, 27, 28, 29, 30,

31, 32, 33, 34, 35, 36, 37, 38, 39, 40,

41, 42, 43, 44, 45, 46, 47, 48, 49, 50

}};

// inc/LowAlarmSet.h

std::array<int, 11> m_lowalarmArray{{

0, 1, 2, 3, 4, 5, 6, 7, 8, 9, 10

}};

// inc/CalvalueSet.h

std::array<int, 5> m_calvalueArray{{ 1, 2, 5, 10, 25 }};

## 4  Step-by-Step Swap Procedure

Step 1 - Hardware

Replace the sensor module. Confirm the transimpedance amp (TIA) gain resistor is correct for the new sensor output range. If TIA output at full scale is under ~1 V, use GAIN_FOUR. If it reaches 1-2 V, use GAIN_TWO.

Step 2 - SensorConfig.h

Comment out the old profile #include. Uncomment the new profile. This automatically changes the BT name, ADS gain, and GasManager defaults.

Step 3 - Four backing arrays

Open SENSOR_CONFIGS.md (or this document, Section 3). Paste the ready-to-paste arrays into RangeSet.h, AlarmSet.h, LowAlarmSet.h, and CalvalueSet.h.

Step 4 - Four menu sections in the .ino

Update: gas list (addGas calls), LIBRARY menu (GasMenuItem calls), Range menu (RangeMenuItem calls), and the inline alarm/calvalue arrays.

Step 5 - Flash

Compile and flash to the device.

Step 6 - Old EEPROM data

The previous sensor's intercept/slope are still in EEPROM. Running a fresh calibration in Step 7 overwrites them automatically. If readings look wrong on first boot, this is expected - calibrate immediately.

Step 7 - Zero calibration

Navigate to CALIBRATION Zero. Expose to zero gas (clean air for CO; zero air cylinder for H2S). Wait 30-60 s for stability. Press S. Saves intercept to EEPROM address 8.

Step 8 - Set Calvalue

Navigate to Calvalue menu. Select the concentration of your certified cal gas cylinder.

Step 9 - Span calibration

Navigate to CALIBRATION Calgas. Apply certified cal gas at 200-500 mL/min. EC sensors stabilize in 60-120 s (longer than PID). Press S when stable. Slope is computed and saved to EEPROM address 16.

Step 10 - Verify

Return to Run screen. Re-apply cal gas - reading should be within +/-5% of certified value. Apply zero gas - should return to 0.

## 5  What Changes vs What Stays the Same

| Item | Changes? | How |
| --- | --- | --- |
| BT name | Yes - automatic | SensorConfig.h |
| ADS gain | Yes - automatic | SensorConfig.h |
| GasManager defaults | Yes - automatic | SensorConfig.h |
| Range / alarm arrays | Yes - manual | Paste from this doc |
| Menu items in .ino | Yes - manual | 4 menu sections |
| Calibration data | Yes - overwritten | Run Zero + Span cal |
| RTC, logging, User ID | No change | - |
| Display, buttons, sleep | No change | - |

## 6  EEPROM Address Map  (GasManager Parameters)

| Address | Content | Size |
| --- | --- | --- |
| 4 | Gas slot index (selected gas) | 4 bytes (int) |
| 8 | Intercept slot 1 | 8 bytes (double) |
| 16 | Slope slot 1 | 8 bytes (double) |
| 24 | Intercept slot 2 | 8 bytes (double) |
| 32 | Slope slot 2 | 8 bytes (double) |
| 40 | Intercept slot 3 | 8 bytes (double) |
| 48 | Slope slot 3 | 8 bytes (double) |
| 56 | Intercept slot 4 | 8 bytes (double) |
| 64 | Slope slot 4 | 8 bytes (double) |
| 72 | Range index | 4 bytes (int) |
| 76 | Alarm index | 4 bytes (int) |
| 80 | Calvalue index | 4 bytes (int) |
| 92 | Low alarm index | 4 bytes (int) |
| 96 | Logging mode | 4 bytes (int) |
| 120 | Cal date string (ASCII) | 31 bytes |
| 210 | Manual log counter | 2 bytes (uint16) |
| 220 | RTC hour | 4 bytes (int) |
| 230 | RTC minute | 4 bytes (int) |

Note: When swapping sensor types with very different sensitivity, the old EEPROM intercept/slope (addresses 8-64) will be overwritten on the first field calibration. You do not need to erase EEPROM manually - just run Zero then Span calibration.